## Slide 1
Yujie Dong
Committee Members:
Dr. Adam W. Hoover, Chair
Dr. Stanley T. Birchfield
Dr. John N. Gowdy
Dr. Eric R. Muth
March 2012
Tracking Wrist Motion to Detect and Measure Eating Intake of Free-living Humans

## Slide 2
Presentation Outline
Motivation
Detecting eating
  Proc. of Int’l. Conf. on Embedded Systems and Applications
  IEEE Transactions on Biomedical Engineering (under review)
Other Work
  Measuring eating consumption via wrist motion tracking (accepted)
  Slowing bite rate reduces energy intake (published)
  An instrumented cafeteria table for recording eating activity (ongoing)
  Measuring system latency from sensing to actuation (under revision)
Summary

## Slide 3: Motivation

## Slide 4
Motivation
Prevalence of obesity
Statistics in 2008
  Worldwide: 1.5 billion adults were overweight and 500 million adults were obese (WHO 2011).
  USA: 34.4% of Americans were overweight and 33.9% of Americans were obese (Flegal et al. 2010).
Related diseases
  Diabetes, heart disease, high blood pressure, stroke, and higher rates of certain cancers (Wellman et al. 2002)
Annual medical cost in the United States
  $147 billion in 2008 (Finkelstein et al. 2009)
  $78.5 billion in 1998

## Slide 5
New tools needed
Joint NSF/NIH Workshop
  (Ershow et al. 2007) “Engineering Approaches to Energy Balance and Obesity: Opportunitites for Novel Collaborations and Research”
Journal of the American Dietetic Association
  (Thompson et al. 2010) “Need for Technological Innovation in Dietary Assessment”
  (McCabe-Sellers et al. 2010) “Advancing the Art and Science of Dietary Assessment through Technology”
  International Obesity Society: 10,000 researchers
  Obesity Society (U.S.): 2,500 researchers

## Slide 6
New tools needed for …
Researchers
Individuals
  Genetics
  Epidemiology
  Clinical outcomes
  Weight loss
  Fitness
  Weight maintenance

## Slide 7
Our vision
Wrist-worn device to passively monitor eating 24/7

## Slide 8: Detecting the eating activities of a free living human using wrist-worn accelerometers and gyroscopes

## Slide 9
Detecting when people are eating
?
?
?

## Slide 10
Related work
Activity recognition using accelerometers
  Walking, running, sitting, and resting (Atallah et al. 2009, Ermes et al. 2008, Karantonis et al. 2006, Khan et al. 2010, Pawar et al. 2007)
  Accidental falls (Yin et al. 2008)
  Sports activities (Parkka et al. 2010)
  Assembly tasks (Ward et al. 2006)
  Activity recognition using gyroscopes
    Walking, sitting, and resting (Ghasemzadeh et al. 2010)
    Tremors associated with Parkinson’s disease (Salarian et al. 2007)
Eating activity recognition? None.
10x power

## Slide 11
Related work
Camera-based approach
  Take pictures before and after eating (Martin et al. 2009)
  Expert compares to database (Williamson et al. 2003)
  Attempts to automate (Zhu et al. 2008)

## Slide 12
Related work
Wearable sensor-based approach
  Throat and ear (Sazonov et al. 2010)
  Arms and back (Amft et al. 2008)

## Slide 13
Our approach
lab testing …… development …….. manufacturing
Wrist-worn sensor
  Increased compliance and social acceptance
  Problems: pattern recognition and power consumption

## Slide 14
Overview of algorithm
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification

## Slide 15
Overview of algorithm
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification

## Slide 16
Raw data
minute
yaw
pitch
roll
x
y
z
Notes: This table shows the notation used in the rest presentation. We are interested in both accelerometer data and gyroscope data, and we are also interested in comparing the performance of different axes. For example, 𝐺_(𝛼,𝑡), 𝐺_(β,𝑡)
, 𝐺_(γ,𝑡) is the gyroscope data along the yaw, pitch and roll axes at time t. 𝐺_(any,𝑡) means gyroscope data along any axis at time t. 𝐺_(sum,𝑡) is the average sum of absolute motion data over a constant window at time t. It can be calculated using this equation.

## Slide 17
Notation:
Preprocessing
device | axis | notation
gyroscope | yaw | 
gyroscope | pitch | 
gyroscope | roll | 
gyroscope | any | 
gyroscope | sum | 
accelerometer | x | 
accelerometer | y | 
accelerometer | z | 
accelerometer | any | 
accelerometer | sum | 
device | axis | notation
gyroscope | yaw | 
gyroscope | pitch | 
gyroscope | roll | 
gyroscope | any | 
gyroscope | sum | 
accelerometer | x | 
accelerometer | y | 
accelerometer | z | 
accelerometer | any | 
accelerometer | sum | 
Linear motion
Rotational motion
Notes: This table shows the notation used in the rest presentation. We are interested in both accelerometer data and gyroscope data, and we are also interested in comparing the performance of different axes. For example, 𝐺_(𝛼,𝑡), 𝐺_(β,𝑡)
, 𝐺_(γ,𝑡) is the gyroscope data along the yaw, pitch and roll axes at time t. 𝐺_(any,𝑡) means gyroscope data along any axis at time t. 𝐺_(sum,𝑡) is the average sum of absolute motion data over a constant window at time t. It can be calculated using this equation.

## Slide 18
Overview of algorithm
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification

## Slide 19
Motion pattern related to eating
Beginning: large wrist motion energy
During eating: reduced energy
End: large wrist motion energy
eat
eat
Example 1
Example 2
Notes: We have found a general pattern of wrist motion energy related to eating activities. At the beginning of an eating activity, there tends to be a period of larger wrist motion energy, probably caused by things like bringing food to a table, adjusting the position of utensils, opening food containers, and unwrapping food. During eating, the total wrist motion energy is reduced. At the end of an eating activity, there tends to be another period of larger wrist motion energy, likely caused by things like putting remaining food away, washing hands, standing up, and putting dishes away.
We investigated the use of both gyroscopes and accelerometers for detecting hypothesized eating activates. We use 𝐺_(sum,𝑡) and 𝐴_(sum,𝑡) to represent the level of wrist motion energy, where a higher value indicates more energetic activity.
The left figure shows an example from our collected data. The dotted lines indicate the start and stop times of the meal in the manual log. It can be seen that the wrist motion energy spikes both before and after the meal.
The right figure shows a second example. In this case, the spike is less pronounced at the end, indicating that the intensity of meal clean up activities was less than during meal preparation.

## Slide 20
Criteria of first eating pattern
Notes: Formally, we detect a hypothesized eating activity based upon five events as illustrated in the figure.
At the beginning of eating, 𝐺_(sum,𝑡) or 𝐴_(sum,𝑡) must surpass a threshold K at time a
The time between a and b should be no more than 10 minutes
At the end of eating, 𝐺_(sum,𝑡) or 𝐴_(sum,𝑡) must surpass the threshold K at time c
The time between b and c should be at least 3 minutes and less than 30 minutes
The time between c and d should be no more than 10 minutes.
If those five conditions are met, then a hpothesized eating activity is detected as having occurred between time b and c

## Slide 21
Criteria of second eating pattern
Notes: If the first two criteria are met , but the data does not meet the third to fifth criteria, we search for a second pattern. It is shown in this figure.
After we find the first and second criteria, we calculate a running average of wrist motion energy starting two minutes after b. At time c, if the following two event are met:
The time between b and c should be at least 3 minutes and less than 30 minutes.
𝐺_(sum,𝑡) or 𝐴_(sum,𝑡) at c is more than 2.5 times the running average.

## Slide 22
Evaluation
|a| < 15 minutes
|b| < 15 minutes
|a| + |b| < 20 minutes

## Slide 23
Overview of algorithm
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification

## Slide 24
Features during hypothesized eating
Mean absolute deviation of yaw, pitch and roll
Regularity of motion
Eating activity
Non-eating activity
Notes: For each hypothesized eating activity, we investigated the use of both gyroscope and accelerometer features for classification. In my presentation, I use gyroscope derived features as an example. The accelerometer features are similar. Here are the most two important types of gyroscope derived features. The first feature is the mean absolute deviation of yaw, pitch, and roll. We assume that this value must be in some medium range, because the wrist motion must vary during eating, but it probably does not vary as much as during other activities.
The second feature is the regularity of motion. We assume an eating activity may have more regular cycles of rest and motion, while other activities may not follow this pattern. This is illustrated in this Figure, where we show two hypothesized activities. The x axis is the time and the y axis is |𝐺_(any,𝑡)|. We define active motion time as the time that rotational motion exceeds the threshold (10deg/s), plus an additional 8 seconds, a typical time between bites we have observed in our previous research. The green rectangles below each activity are the active motion time. Thus we define the regularity of motion as the total active motion time over the total activity time. For the top activity, it can be seen that the regularity of motion is much larger than for the bottom activity, even though the total time that motion exceeds the threshold K is the same.

## Slide 25
Features during hypothesized eating
Line fit (slope of energy over time)
Entropy
Variance
Zero crossing rate (frequency of motion)
Mean of minima (size of motion)
Mean of maxima
Percentile distribution (range of motion)
Auto correlation (repetition of motion)
Number of bites
Mean of bite interval
Variance of bite interval
Skewness of bite interval

## Slide 26
Overview of algorithm
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification

## Slide 27
Classification: notation
Class space: V = {m, n}
  m: actual eating activity
  n: non-eating activity
Feature space: A = {}
  : the value of feature set i in hypothesized eating j
Notes: Before I discuss the classification methods, I will introduce the notation first.
Our class space contains two sets, m means actual eating activity and n means non-eating activity. Look at this figure, all the hypothesized detected eating activities were manually grouped into either actual eating activities or non-eating activities, using
  Our feature space contain N features, namely 𝑎_1,𝑎_2,...,𝑎_𝑁. For example, a1 means mean absolute deviation of yaw, a2 means mean absolute deviation of pitch, and so on. the The notation 𝑎_(i,j) means the value of feature set i in hypothesized eating j. For example, a11 means mean absolute deviation of yaw in the first hypothesized eating. a12 means mean absolute deviation of yaw in the second hypothesized eating.

## Slide 28
Classification: FLD
For each feature, calculate
Fisher linear discriminant (FLD): J=
A higher FLD indicates the feature has better discriminatory power (is better for classification)
mean of feature for actual eating
variance of feature for actual eating
mean of feature for non-eating
variance of feature for non-eating
Notes:   We use the fisher linear discriminant (FLD) to determine a rough measure of the discriminatory power of each feature. For each feature described, such as mean absolute deviation of yaw, regularity of yaw, we calculated 𝜇_𝑚 and 〖𝜎_𝑚^ 〗^2 as the mean and variance of feature for actual eating, and 𝜇_𝑛 and 〖𝜎_𝑛^ 〗^2 as the mean and variance of feature for non-eating activities. The FLD was then calculated as this equation.
  A higher FLD indicates the feature has better discriminatory power which is better for classification.

## Slide 29
Classification: Naïve Bayes
Using Bayes theorem
: mean of feature associated with class
: variance of feature associated with class
Assume attributes are conditionally independent
classification
eating or other
features
Notes: We selected sets of features having higher FLD values and used a Naïve Bayes classifier. The Bayesian approach is to use feature values a1, …, an to assign the hypothesized eating which is actual eating activities or non-eating activities to the most probable class 𝑣_𝑀𝐴𝑃.
Using Bayes theorem, we can rewrite this expression as this equation.
In practice, we are only interested in the numerator of that fraction since the denominator does not depend on 𝑣_𝑗. The Naïve Bayes classifier is based on the simplifying assumption that the attribute values are conditionally independent. Thus the model can be expressed in this equation.
In this equation, ├ 𝑃(𝑣_𝑗 ) is the risk factor, ├ 𝑃(𝑎_(𝑖,𝑗) |𝑣_𝑗 ) is computed by plugging a_ij into the equation for a normal Gaussian distribution.

## Slide 30
Evaluation
Five folder cross validation scheme (Kohavi 1995)
  Four-fifths for training, one-fifth for testing
  Repeat 5 times
  Average the accuracy

## Slide 31: Data Collection

## Slide 32
Data collection: hardware
iPhone 4
  3 axis accelerometers
  3 axis gyroscopes
  Large memory (16 GB)
  Large battery (1420 mAh)

## Slide 33
Data collection: recording program
Record 6 axes at 60 Hz
Store data on the device
Transfer to a computer through USB
Current time and status
Start, stop, and reset buttons

## Slide 34
Data collection: protocol
Start recording
Conduct activity
Stop recording
Remove device
Keep food log

## Slide 35
Data collection: manual log entry

## Slide 36
Data collection: participants
30 participants
Approved by Clemson University Institutional Review Board
Age: 18 ~ 32
Gender
  12 male
  18 female
Ethnicity
  25 Caucasian
  2 Asian or Pacific Islander
  2 African American
  1 Hispanic

## Slide 37
Usable data
meals
device on

## Slide 38
Usable data
Category | Occurrences
Total manually logged eating activities | 78
Usable eating activities | 35
Subject forgot to manually log start or stop time | 4
Motion data was not recorded during manually logged entry | 8
Activity duration is too short (less than 3 minutes) | 3
Activity duration is too long (more than 30 minutes) | 9
Subject started or stopped recording less than 10 minutes before or after eating | 18
Logged eating activity is inside another logged eating activity | 1
American Time Use Survey: 1.25 hours eating and drinking per day (Bureau of Labor Statistics 2012)

## Slide 39: Results

## Slide 40
Result
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification
Results

## Slide 41
Detected hypothesized eating activities
 | True detection | False detection | Undetected
Accelerometer | 31 | 203 | 4
Gyroscope | 31 | 267 | 4
Accelerometers and gyroscopes detect the same amount of actual eating activities
Gyroscopes have more false detections

## Slide 42
Result
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification
Results

## Slide 43
FLD of accelerometer features
Accelerometer feature | FLD | TP | FP
mean absolute deviation of x | 0.05 | 21 | 111
mean absolute deviation of y | 0.00 | 26 | 133
mean absolute deviation of z | 0.00 | 24 | 117
regularity of x motion | 0.26 | 28 | 124
regularity of y motion | 0.26 | 28 | 112
regularity of z motion | 0.28 | 27 | 109
accelerometer line fit | 0.01 | 13 | 86

## Slide 44
FLD of gyroscope features
Gyroscope feature | FLD | TP | FP
mean absolute deviation of yaw | 0.44 | 22 | 63
mean absolute deviation of pitch | 0.43 | 23 | 61
mean absolute deviation of roll | 0.78 | 22 | 43
regularity of yaw motion | 1.04 | 26 | 53
regularity of pitch motion | 0.80 | 25 | 52
regularity of roll motion | 1.08 | 24 | 55
gyroscope line fit. | 0.00 | 18 | 80
bite count | 0.74 | 25 | 58
mean of bite interval | 0.16 | 23 | 74
variance of bite interval | 0.01 | 24 | 93
skewness of bite interval | 0.32 | 16 | 50

## Slide 45
Result
Preprocessing
Detect hypothesized eating
Feature extraction
Final classification
Results

## Slide 46
Naïve Bayes classification
Gyroscope feature | Total feature | True detection | False detection | Undetected
reg yaw, reg roll | 2 | 25 | 54 | 6
mad roll, reg roll | 2 | 26 | 46 | 5
reg yaw, reg pitch, reg roll | 3 | 27 | 54 | 4
mad roll, reg roll, bite count | 3 | 27 | 47 | 4
mad roll, reg yaw, reg pitch, reg roll | 4 | 27 | 51 | 4
mad roll, var roll, reg roll, bite count | 4 | 24 | 43 | 7
mad roll, reg yaw, reg pitch, reg roll, bite count | 5 | 27 | 49 | 4
mad roll, var roll, reg roll, line fit, bite count | 5 | 23 | 43 | 8
2 features are enough
wrist roll gives best rate

## Slide 47
Conclusion
Three-part pattern
  pre-eating
  during-eating
  post-eating
Detection: accelerometers
Classification: gyroscopes
Pilot test identified 26 of 35 (74%) eating activities with 1.3 false positives per actual eating activity

## Slide 48: Other Work

## Slide 49: Monitoring eating consumption via wrist motion tracking

## Slide 50
“Bite counter” device
Worn like a watch
Automatically measure bites
Used in natural environment
Feedback in real time

## Slide 51
“Bite counter” device

## Slide 52
Our invention
High correlation between a simple pattern of wrist roll and the taking of a bite

## Slide 53
Result
Exp 1: Test across a large number of subjects
  51 subjects
  139 meals
  Eat Kellogg’s Eggo cinnamon toast waffles
  Use fork
  94% sensitivity
  80% PPV
Exp 2: Test across a large variety of foods
  47 subjects
  49 meals
  Eat their own foods
  Use their own utensils
  86% sensitivity
  81% PPV

## Slide 54
Result
Exp 3: Test relationship between bites and calories
  4 subjects
  54 meals
  Eat their own foods
  Use their own utensils
  Correlation: 0.6

## Slide 55
Result
Exp 4: Test relationship between bites and calories
  83 subjects
  3246 meals
  76% have correlation > 0.4
  69% between 0.4 and 0.7
Correlation: 0.4
Correlation: 0.7

## Slide 56: Slowing bite rate reduces energy intake

## Slide 57
Slowing bite rate reduces intake
30 subjects, each ate three meals
  Baseline
  Bite-rate feedback
  50% Slower bite-rate target
Slower bite-rate condition significantly affect participants who ate more than 400 kcal at baseline, but not the rest.

## Slide 58: An instrumented cafeteria table for recording eating activity

## Slide 59
Cafeteria eating research
A dining table equipped with video recorder, scale, and wrist-worn motion sensor
Conduct food intake measurement in a cafeteria setting so calories and actual bites can be measured
Dining table
Recording program

## Slide 60: Measuring system latency from sensing to actuation

## Slide 61
System latency measurement
Two systems
  A camera for sensing, an LCD for actuation
  An orientation sensor for sensing, a motor for actuation
Our finding
  A cyclical variation in latency
  Orientation sensor error looks like a variation in latency

## Slide 62
Summary
Main contributions
  Develop an automated method to detect eating activities by tracking wrist motion throughout the day
  Develop an automated method to measure eating intake by tracking wrist motion during a meal
Other contributions
  Find slowing bite rate can reduce energy intake
  Explore the relationship between bites and calories
  Explain the variation in system latency measurement

## Slide 63
Publications (Journal)
Y. Dong, A. Hoover, J. Scisco, and E. Muth, “Detecting the eating activities of a free living human using wrist-worn accelerometers and gyroscopes”, in IEEE Transaction on Biomedical Engineering, 2012 (under review).
Y. Dong, A. Hoover, J. Scisco, and E. Muth, “Monitoring eating consumption via wrist motion tracking”, in Applied Psychophysiology and Biofeedback, 2011 (accepted).
W. Wu, Y. Dong, and A. Hoover, “Measuring digital system latency from sensing to actuation at continuous 1 ms resolution”, in Presence: Teleoperators and Virtual Environments, 2011 (under revision).
J. Scisco, E. Muth, Y. Dong, A. Hoover, “Slowing bite-rate reduces caloric consumption; an application of the bite counter device”, in Journal of the American Dietetic Association, vol. 111, 2011, pp. 1231-1235.

## Slide 64
Publications (Conference)
Y. Dong, A. Hoover, J. Scisco, and E. Muth, “Detecting eating using a wrist mounted device during normal daily activities”, Proceedings of the 9th International Conference on Embedded Systems and Applications, 2011.
J. Scisco, E. Muth, Y. Dong, A. Hoover, P. O'Neil, and S. Fishel-Brown, “Usability and acceptability of the bite counter device”, Proceedings of the Human Factors and Ergonomics Society Annual Meeting, 2011.
W. Wu, Y. Dong, and A. Hoover, “Measuring digital system latency from sensing to actuation at continuous 1 ms resolution” [abstract], The 3rd International Symposium on Visual Image Safety, 2011.
J. Scisco, E. Muth, A. Hoover, Y. Dong, L. Hill, M. Wiles, and S. Harris, “An environmental cue to eat slowly reduces food consumption” [abstract], Proceedings of the 55th annual meeting of the Human Factors and Ergonomics Society, 2010.
Y. Dong, A. Hoover and E. Muth “A device for detecting and counting bites of food taken by a person during eating”, Proceedings of IEEE Conference on Bioinformatics and Biomedicine, 2009, pp. 265-268.

## Slide 65
Publications (Patent)
A. Hoover, E. Muth, Y. Dong and J. Scisco, “Eating Activity Detection Device”, disclosed to Clemson University for patent application in March, 2012
A. Hoover, E. Muth and Y. Dong, “Weight control device”, US Provisional Patent Application serial no. 61/144.203, filed January 13, 2009