## Slide 1: Reverse-Projection Method for Measuring Camera MTF
Stan Birchfield

## Slide 2: MTF measures camera sharpness
Optical transfer function (OTF) Fourier transform of impulse response
Modulation transfer function (MTF) Magnitude of OTF
Spatial frequency response (SFR) MTF for digital imaging systems
Slanted edge Simplest target for computing SFR
High MTF Low MTF  Sharp camera  Blurry camera
SFR ≈ MTF

## Slide 3: What affects MTF?
MTF affected by
Lens
Sensor
Electronics
Target
Pose
Lighting intensity and color
Color channel
ROI
Orientation
Algorithm


Slanted edge

## Slide 4: Variability of results
Same:
Sensor
Target
Lighting conditions
Pose
Different:
Crops from 25 to 50 pixels wide
Angles between 5 and 20 degrees
Set of ~20 total cropped images
MTF30

## Slide 5: Variability of results
Same:
Sensor
Target
Lighting conditions
Pose
Different:
Crops from 25 to 50 pixels wide
Angles between 5 and 20 degrees
Set of ~20 total cropped images

## Slide 6: Variability of results
These are all crops of the same image!
Same:
Sensor
Target
Lighting conditions
Pose
Angle
Different:
Crops from 25 to 50 pixels wide

## Slide 7: ISO 12233: The standard way to estimate MTF
Define ROI (region of interest) either manual or automatic
Gamma expansion undo OECF (opto-electronic conversion function)
Luminance from RGB
Find points along edge
Fit line to points
Project 2D array onto 1D array
Differentiate
Apply Hamming window
Compute DFT (discrete Fourier transform)
Magnitude of DFT yields estimate of MTF
[2000, 2014]
Slanted edge approach

## Slide 8: ISO 12233: The standard way to estimate MTF
Define ROI (region of interest) either manual or automatic
Gamma expansion undo OECF (opto-electronic conversion function)
Luminance from RGB
Find points along edge
Fit line to points
Project 2D array onto 1D array
Differentiate
Apply Hamming window
Compute DFT (discrete Fourier transform)
Magnitude of DFT yields estimate of MTF
[2000, 2014]
Slanted edge approach

## Slide 9: Comparison
Standard approach
Centroid of derivative along rows
Ordinary least squares
Forward projection
Proposed approach
Ridler-Calvard segmentation
Total least squares
Reverse projection
Step 4: Find points along edge
vs.
Step 5: Fit line to points
vs.
Step 6: Project from 2D to 1D
vs.

## Slide 10: Step 4: Find points along edge
Derivative along rows
For each row:
  Compute 1D derivative using FIR filter
  Compute centroid
Local
Assumes near-vertical orientation
Ridler-Calvard segmentation
Repeat:
  m1  graylevel mean below t
  m2  graylevel mean above t
  t  (m1 + m2) /2
Global
Any orientation
Standard approach
Proposed approach
T. W. Ridler, S. Calvard, Picture thresholding using an iterative selection method, IEEE Trans. System, Man and Cybernetics, SMC-8:630-632, 1978.

## Slide 11: Step 5: Fit line to points
Ordinary least squares
y = mx + b (slope-intercept form)
Minimizes vertical error
Affected by orientation
Total least squares
ax + by + c = 0 (standard form)
Minimizes perpendicular error
Any orientation
Standard approach
Proposed approach

## Slide 12: Step 6: Project from 2D to 1D
Forward projection
For each 2D pixel:
  Project to nearest 1D bin
  Then average
Reverse projection
For each 1D bin:
  Project average of 2D pixels (maybe interpolated)
Standard approach
Proposed approach
all
bins
have
values!
some
bins
may be
empty

## Slide 13: Forward vs. reverse projection
Forward projection
Reverse projection
(All other steps identical)
bump

## Slide 14: What causes the bump?
Forward projection
Reverse projection
period of noise = amount of oversampling

## Slide 15: Stability across subsamplings
more stable

## Slide 16: Experiment 1: Different orientations
0o to 45o, increments of 5o (50x50 crops)
Details: 4:1 contrast ratio, 580 lux, 5540 K, AGC off, JPEG
gap

## Slide 17: MTF curves of reverse projection
consistency
across
orientation

## Slide 18: Experiment 2: Vertical crop
Number of rows: 3 to 50

## Slide 19: Experiment 3: Horizontal crop
Number of columns: 5 to 50

## Slide 20: Experiment 4: Square crop
Side length: 5 to 50

## Slide 21: MTF curve comparison
50 x 50 crop
5o orientation
 interpolation alone does
not explain discrepancy
Bilinear interpolation
Bicubic interpolation
No interpolation

## Slide 22: Conclusion
Existing implementations of ISO 12233 exhibit variability
  ± 0.07 in MTF30 due to crop alone
  significant digits
Proposed solution:
  Ridler-Calvard to find points
  total least squares to fit line
  reverse projection
Advantages:
  stability
  invariance to orientation
Remaining questions:
  Is high-frequency bump caused by forward projection?
  Is reverse projection more or less accurate?
  More experiments?

## Slide 23: Thanks!
Special thanks to
Andrew Juenger
Erik Garcia
Kanzah Qasim